# OpenGammaKit

Assembly Guide v0.1.2

Exploded Assembly View

Parts List

- Cable gland

- Enclosure tube

- Main board

- SiPM board

- Spacer ring

- Scintillation crystal NaI(Tl)

- End cap

Assembly Sequence

- Insert the USB cable (3.5 mm diameter) through the cable gland (1).

- Solder the USB Micro connector to the cable, adding a 4.7 – 5.1 kΩ resistor between the CC (Configuration Channel) and GND pins (see Appendix A).

- Feed the cable with the connector through the enclosure tube (2).

- Secure the cable gland (1) inside the enclosure to hold the cable firmly in place.

- Solder the SiPM board (4) to the main board (3) and connect it to the USB Micro connector.

- Insert the main board (3) into the internal guide slots of the enclosure tube (2), pulling the cable through the cable gland.

  - The SiPM board should rest precisely on the inner support ring inside the enclosure.

  - The SiPM sensor must face outward from the enclosure.

- Tighten the cable gland on the outside of the enclosure.

- Apply a small drop of silicone oil to the surface of the SiPM to ensure good optical coupling.

- Insert the spacer ring (5) into the upper part of the enclosure.

- Place the NaI(Tl) scintillation crystal (6) inside the spacer ring.

- Screw the end cap (7) onto the top of the enclosure with light tension—tight enough to secure the crystal without applying excessive force.

- (Optional – Light Shielding):
If the enclosure is made of optically transparent plastic, the SiPM must be protected from external light exposure.

  - Slide a black heat-shrink tube (50 mm diameter) over the assembled enclosure.

  - Gently heat the shrink tubing to fit snugly, taking care not to overheat or deform the plastic enclosure.

Notes

- Ensure that all solder joints are clean and free of flux residues.

- Avoid touching the surface of the SiPM or the crystal directly—use gloves or optical tissue if necessary.

- Verify correct alignment before sealing the enclosure.

Appendix A

USB Type-C Cable schema